## Real time Fatigue
Real time fatigue of porcine aortic heart valve roots
Fig 7 Valve Fully Closed Images (where LC: left coronary leaflet, RC: right coronary leaflet and NC: non-coronary leaflet) of the (a) Cellular (b) Decellularised Porcine Aortic Heart Valve Roots at Maximum 1.2 Million Cycles with Leaflet Tears Circled in Red, Captured During Real Time Fatigue Assessment
Cellular
Decellularised
(a)
(b)
NC
RC
LC
RC
LC
NC
Failure
LC
RC
NC
LC
RC
NC
Failure
LC
RC
NC
LC
NC
LC
RC
NC
Failure
LC
RC
NC
LC
RC
NC
LC
RC
NC
RC
Failure
RC
LC
C1
C2
C3
C4
C5
C6
D1
D2
D3
D4
D5
D6
Valve ID Number
Valve ID Number
Valve ID Number

## Competency
Cellular Aortic Heart Valve Root Aortic Heart Valve Roots Aortic Heart Valve Roots |  |  |  |  | Decellularised Aortic Heart Valve Root Aortic Heart Valve Roots Aortic Heart Valve Roots
Leakage Flow Rate
| Valve ID Number | Pre-test | Mid-test | Post-test | Valve ID Number | Pre-test | Mid-test | Post-test
|  | ms |  |  |  | ms
| C1 | >1800 | >1800 | 352 | D1 | 749 | xxx | xxx
| C2 | >1800 | >1800 | 1005 | D2 | 1741 | 1550 | >1800
| C3 | >1800 | >1800 | xxx | D3 | 951 | 324 | 255
| C4 | >1800 | >1800 | 1284 | D4 | >1800 | >1800 | >1800
| C5 | >1800 | 92 | xxx | D5 | 685 | 72 | 1245
| C6 | >1800 | 1872 | 154 | D6 | 938 | 1245 | 925
|  | mL.s-1 |  |  |  | mL.s-1
| C1 | <0.85 | <0.85 | 4.358 | D1 | 2.0481 | xxx | xxx
| C2 | <0.85 | <0.85 | 1.5264 | D2 | 0.8811 | 0.98968 | <0.85
| C3 | <0.85 | <0.85 | xxx | D3 | 1.613 | 4.7346 | 6.0157
| C4 | <0.85 | <0.85 | 1.19 | D4 | <0.85 | <0.85 | <0.85
| C5 | <0.85 | 16.6739 | xxx | D5 | 2.2394 | 21.3056 | 1.2321
| C6 | <0.85 | 0.81944 | 9.96 | D6 | 1.6354 | 1.2321 | 1.6584
xxx - not measured
Pre, Mid and Post- Test Mean Static Leakage Flow Rates of Each Cellular and Decellularised Porcine Aortic Heart Valve Root

## Pulsatile Flow
Pulsatile flow performance of porcine aortic heart valve roots
|  | Cellular Porcine Aortic Heart Valve Root |  |  |  | Decellularised Porcine Aortic Heart Valve Root
|  | Heart Rate | Root mean square (RMS) forward flow (QRMS) | pressure difference during forward flow (ΔP) | Effective orifice area (EOA) | Heart Rate | Root mean square (RMS) forward flow (QRMS) | pressure difference during forward flow (ΔP) | Effective orifice area (EOA)
|  | bpm | mL/s-1 | mmHg | cm2 | bpm | mL/s-1 | mmHg | cm2
|  |  | 0 | 0 |  |  | 0 | 0
|  | 60 | 149.0971 | 2.616 | 2.0765 | 60 | 169.5933 | 2.4017 | 2.092
|  | 72 | 213.6642 | 3.563 | 2.3285 | 72 | 227.8966 | 4.9753 | 1.9784
|  | 80 | 224.7272 | 3.9556 | 2.1504 | 80 | 260.1841 | 6.6151 | 2.0443
|  | 100 | 343.214 | 8.0995 | 2.347 | 100 | 368.7011 | 11.8528 | 2.1162
Fig 9 Valve Fully Closed Images of (a) Cellular (b) Decellularised Porcine Aortic Heart Valve Roots at Heart Rate 72 bpm during Pulsatile Flow Assessment, captured from the Front View by a High-Speed Camera with Leaflet Tears Circled in Red
NC
Cellular
Decellularised
(a)
(b)
RC
LC
RC
LC
NC
LC
RC
LC
NC
LC
NC
LC
RC
NC
LC
RC
LC
RC
NC
RC
Failure
RC
LC
C1
C3
C4
C6
D1
D2
D3
D4
D5
D6
Valve ID Number
Valve ID Number
Valve ID Number
LC
RC
NC
NC
Failure
NC
Failure
RC

## Leaflet Properties
Uniaxial tensile properties of aortic leaflets
Radial
| Ec (MPa) Cellular | Ec (MPa) Decellularised | E e (MPa) Cellular | E e (MPa) Decellularised | σ UTS (MPa) Cellular | σ UTS (MPa) Decellularised
| 1.6593 | 2.735 | 0.0072 | 0.0309 | 0.7155 | 1.3659
| 0.5482 | 1.859 | 0.019 | 0.0281 | 0.3319 | 0.5
| 0.3082 | 1.3834 | 0.0257 | 0.0205 | 0.2183 | 0.6424
| 1.2111 | 2.5663 | 0.0112 | 0.0297 | 0.5576 | 1.1851
|  | 5.2589 |  | 0.0182 |  | 2.1569
|  | 0.9861 |  |  |  | 0.5333
Mean ± 95 % Condidence Limits | 0.93 ± 0.98 | 2.47 ± 1.60 | 0.02 ± 0.01 | 0.03 ± 0.01 | 0.46 ± 0.36 | 1.06 ± 0.68
Circumferential
| 7.8837 | 38.115 | 0.0133 | 0.426 | 2.2018 | 8.2963
| 7.878 | 49.815 | 0.03 | 0.066 | 2.94 | 10.051
| 3.287 | 55.109 | 0.0138 | 0.1799 | 2.0649 | 8.825
| 15.315 | 60.505 | 0.0213 | 0.3152 | 2.713 | 12.034
|  | 61.524 |  | 0.5678 |  | 9.0583
|  | 54.442 |  |  |  | 7.5815
Mean ± 95 % Condidence Limits | 8.59 ± 7.92 | 53.44 ± 8.99* | 0.02 ± 0.01 | 0.31 ± 0.25* | 2.48 ± 0.66 | 9.31 ± 1.64*
Porcine Aortic Heart Valve Roots (Mean Value Table)
Mean Tensile Parameters for Leaflets | Cellular |  |  | Decellularised |  | p value
| N |  |  | N
Radial
Ec (MPa) | 4 |  | 0.93 ± 0.98 | 6 | 2.47 ± 1.60 | 0.1
E e (MPa) | 4 |  | 0.02 ± 0.01 | 5 | 0.03 ± 0.01 | 0.07
σ UTS (MPa) | 4 |  | 0.46 ± 0.36 | 6 | 1.06 ± 0.68 | 0.11
Circumferential
Ec (MPa) | 4 |  | 8.59 ± 7.92 | 6 | 53.44 ± 8.99* | 1e-05
E e (MPa) | 4 |  | 0.02 ± 0.01 | 5 | 0.31 ± 0.25* | 0.02
σ UTS (MPa) | 4 |  | 2.48 ± 0.66 | 6 | 9.31 ± 1.64* | 3e-05